## Slide 1: Edwards Alterra Adaptive Prestent with the SAPIEN 3 Transcatheter Pulmonic Valve Delivery System
See instructions for use for full prescribing information, including indications, contraindications, warnings, precautions and potential adverse events
Edwards Confidential. Do not reproduce or distribute.
DOC-0555472 Rev B
Pulmonic Procedural Training Manual

## Slide 2: Table of contents
04
Procedural animation
Alterra adaptive prestent: Device specifications
06
SAPIEN 3 Pulmonic delivery system: Device specifications
13
Patient screening and preparation
26
Alterra procedural overview
45
Alterra troubleshooting
75
SAPIEN 3 Pulmonic delivery system procedural overview
81
Intra-procedural considerations
103
Complications
105

## Slide 3: Procedure overview
Valve delivery
Alterra delivery
Final implant
The procedure and system is designed as a two phase, fully percutaneous approach
Alterra
SAPIEN 3 valve
Alterra delivery system
Pulmonic delivery system
1
2

## Slide 4: Procedural animation

## Slide 5: Alterra adaptive prestent with the SAPIEN 3 Pulmonic delivery system
Click above to pause/play video

## Slide 6: Alterra adaptive prestent: Device specifications

## Slide 7: Edwards Alterra adaptive prestent system
Outflow
Suture band (29 mm SAPIEN 3 landing zone)
Inflow sealing zone
12 outflow apices
Open cells
12 inflow apices including 2 connector tabs
Waist markers (3)
PET fabric
Nitinol frame

## Slide 8: Alterra adaptive prestent dimensions
19 mm
27 mm
39 mm
41 mm
49 mm
45 mm
45 mm

## Slide 9: Edwards Alterra adaptive prestent delivery system
Guidewire lumen/ flushport
Tapered tip
Retractable outer shaft
Delivery system marker band (not visible without fluoro)
Deployment wheel
Prestent connector tab
Prestent connector

## Slide 10: Edwards Alterra adaptive prestent delivery system
Delivery system dimensions | 
Working length (distance from distal end of tapered tip to handle) | 105 cm
Tapered tip length | 45 mm
Alterra crimped length (apices to apices) | 59 mm
Alterra capsule OD | 6.7 mm / 20Fr
Alterra shaft OD | 5.7 mm / 17Fr
Wire compatibility | 0.035”
Alterra adaptive prestent is compatible with the 16Fr eSheath or equivalent | 
Prestent connector
Alterra waist markers
Radiopaque marker band
Tapered tip
NOTE
    Arterial vascular size for 16Fr eSheath is a minimum vessel diameter of 6 mm.

## Slide 11: Edwards eSheath introducer set
Alterra adaptive prestent is compatible with the 16Fr Edwards eSheath or equivalent
Hydrophilic coating to facilitate easy insertion
57 cm
36 cm
9.5 cm Partially expandable
26.5 cm Fully expandable
See the eSheath IFU for full prescribing information
NOTE
    Arterial vascular size for 16Fr eSheath is a minimum vessel diameter of 6 mm.
WARNING
    It is not recommended to use the long Gore DrySeal sheath with the Alterra delivery system.

## Slide 12: Edwards eSheath introducer set
Duckbill valve
Disc valve
Cross-slit valve
Tri-seal valve technology designed for hemostasis
Seal |  | Hemostasis
 | Duckbill valve | Designed to provide hemostasis when nothing is inside sheath
 | Disc valve | Designed to provide hemostasis with a 5Fr or greater catheter
 | Cross-slit valve | Designed to provide hemostasis for guidewires ≥ 0.035”

## Slide 13: SAPIEN 3 Pulmonic delivery system: Device specifications

## Slide 14: SAPIEN 3 Pulmonic System with the 29 mm SAPIEN 3 Transcatheter heart valve
Edwards SAPIEN 3 Transcatheter heart valve
Crimper
Pulmonic delivery system
Locking Syringe
Color-coded system
Accessories
Orange
29 mm THV

## Slide 15: SAPIEN 3 Transcatheter heart valve
1 No clinical data are available with which to evaluate the long-term impact of the Carpentier-Edwards ThermaFix tissue process in patients
2 Rounded to the nearest 0.5 mm
Bovine pericardial tissue
Scalloped leaflet shape
Utilizes the same bovine pericardial tissue and processes as Edwards surgical valves
Outer sealing skirt
PET outer sealing skirt designed to minimize paravalvular leak
Inner skirt
Polyethylene terephthalate (PET)
SAPIEN 3 THV size | 29 mm
Tissue | Bovine pericardium, Thermafix tissue process1
Frame | Cobalt-chromium alloy
Crimped height2 | 31 mm
Expanded height | 22.5 mm
Foreshortening | 8.5 mm

## Slide 16: SAPIEN 3 Transcatheter heart valve – Frame design
SAPIEN 3 valve
Frame design
Wide strut angles contribute to ultra-low delivery profile
Four rows and columns between commissures for high radial strength
Cobalt chromium alloy frame and wide strut angles provide fatigue resistance and high radial strength for circularity

## Slide 17: SAPIEN 3 Pulmonic delivery system overview
Designed for valve deliverability
Valve capsule to cover SAPIEN 3 valve
Hydrophilic coating on tapered tip and valve capsule
Valve crimped directly on balloon
Inline sheath
Color coded gripper
Tapered tip to facilitate tracking
Valve capsule to cover valve
Balloon folding technology simplifies the procedure
Eliminates valve alignment and flex catheter retraction steps

## Slide 18: SAPIEN 3 Pulmonic delivery system – Distal end
Valve is covered by capsule until valve is positioned in desired location
Valve is crimped directly on balloon (no valve alignment)
Animal model shown for illustration
Internal shoulders (underneath balloon)
Valve capsule
Tapered tip
Balloon shaft
Radiopaque shoulder markers
Tapered tip
Valve capsule

## Slide 19: SAPIEN 3 Pulmonic delivery system – Handle
Color-coded gripper
Delivery system lock
Inflation volume
Outer shaft
Flush port
Balloon shaft
Balloon inflation port
Guidewire lumen
Balloon shaft marker

## Slide 20: SAPIEN 3 Pulmonic delivery system – Inline sheath
10 cm effective length
25.5Fr / 8.5 mm outer diameter (all sizes)
Flush port
NOTE
    The inline sheath does not come off the delivery system. Shown separate for illustrative purposes.

## Slide 21: SAPIEN 3 Pulmonic delivery system – Dilators
40 cm
THV size | Provided dilator size
29 mm | 28Fr

## Slide 22: SAPIEN 3 Pulmonic delivery system – Specifications
THV Size | Catheter capsule outer diameter
29 mm | 28.5Fr / 9.5 mm
112 cm
Effective length balloon uncovered
42 mm Tapered tip
Pulmonic delivery system
Guidewire compatibility
0.035”

## Slide 23: Crimper, Qualcrimp crimping accessory and 2-piece crimp stopper
Additional components used during device preparation
  Crimper is the same for 29 mm THV size and is packaged separately
  Qualcrimp crimping accessory is designed to protect the THV leaflets when crimping to a low profile and is packaged with the delivery system
  Color-coded 2-piece crimp stopper is used during crimping and is packaged with the delivery system
Qualcrimp crimping accessory
Crimper
2-piece crimp stopper
29 mm

## Slide 24: Locking Syringe
For use with 29 mm THV system
Locking syringe is used for deployment of the THV and pre-dilation balloon
Volume-based inflation
Lock/unlock symbols on handle
1 mL graduation marks for volume accuracy
Locking syringe (38 mL)
Used with the 29 mm THV system
Customized long, narrow barrel for inflation with large volumes
Pressure gauge removed for device placement during inflation

## Slide 25: Recommended equipment
Sterile table equipment
6 sterile rinsing basins
  2 – physiological saline (1 liter) – valve rinsing
  2 – physiological saline (1 liter) – Qualcrimp rinsing
  1 – heparinized saline (1 liter) – flushing
  1 – diluted radiopaque contrast medium (15 cc contrast: 85 cc saline)
20 cc syringes or larger
50 cc syringes or larger
High-pressure 3-way stopcock
Prep table
Sterile forceps
Sterile scissors
Room equipment
Vascular dilators
Balloon tip catheter
Angiographic catheters
Exchange length 0.035" (0.89 mm) extra-stiff guidewires: i.e., Lunderquist, Meier, Amplatz extra stiff
Intra-cardiac echocardiographic system (optional)
Vascular closure devices
Large bore short sheath
Temporary pacemaker and pacing leads
Liter heparinized saline
Liter normal 0.9% saline
Standard cardiac catheterization lab equipment
Fluoroscopy (bi-plane, fixed, mobile or semi-mobile fluoroscopy systems appropriate for use in percutaneous coronary interventions)
Sizing balloons: i.e., PTS – 30 mm and 40 mm diameters
Equipment provided by Edwards Lifesciences
Edwards SAPIEN 3 valve
Edwards Pulmonic delivery system
Edwards crimper
Edwards provided inflation device
Edwards eSheath introducer set
Edwards Alterra delivery system

## Slide 26: Patient screening and preparation

## Slide 27: Alterra design
Fixation
The apices are designed to engage with the anatomy and are critical to provide prestent stability
The Alterra waist and shoulders are not designed to be the primary means of fixation
Sealing
Provide sealing at the inflow sealing zone
Sit within the anatomy and provide apposition against the RVOT/PA at both the inflow sealing zone and outflow portion of the Alterra
Open cells
Provide flow to the branch pulmonary arteries when the Alterra is placed at or across pulmonary bifurcation
Waist
Sit in the MPA and provide a landing zone for the SAPIEN 3 valve
Provide valve retention by the suture band embedded in the waist of the Alterra
Valve retention is not dependent on anatomical support
Positioning
Positioning is not dependent on native annular location

## Slide 28: Pre-screening
Used to determine whether to obtain a retrospectively gated CT to send to Edwards for analysis
Measure RVOT length and diameters on MRI or CT
Sizing guidelines (measurements are approximate, as analysis is not based on diameters alone and is performed in systolic phase by CTA)
  RVOT diameters should be approximately 25-40 mm throughout the landing zone
  RVOT length should be approximately ≥48 mm from contractile RV to back wall of MPA
  Dimensions slightly out of this range do not necessarily preclude Alterra candidacy
If patient’s RVOT is within these criteria, obtain a retrospectively gated CT (without dose modulation) and send to Edwards for analysis and positioning
Assess vascular access (patent femoral vein or other venous access required)
Click above to pause/play video

## Slide 29: CT image acquisition
Parameters | 
Scan coverage | Minimum coverage: RVOT, MPA and branch PAs Coverage can be extended to include RV and LV if CT is used for chamber quantification (higher radiation exposure!)
Heart cycle coverage | Systole and diastole should be covered (ideally the entire cardiac cycle), to allow for dynamic assessment of landing zone/potentially contractile RVOT
Data acquisition mode | Volume CT scanner (GE Revolution, Canon/Toshiba AQ1): ECG-gated volume acquisition All other scanner: retrospective ECG-gating
Contrast application | Optimized for right heart/MPA assessment Biphasic injection protocol using diluted contrast (contrast/saline mixture) to reduce streak artifacts
Image reconstruction | Multiphasic reconstruction covering entire cardiac cycle in 5-10% or 50msec increments Maximize use of iterative reconstruction

## Slide 30: CT image acquisition
Consideration in regard to image quality and radiation dose
Higher image noise compared to coronary CTAs is acceptable as RVOT/MPA are large structures
Maximize utilization of iterative image reconstruction
  Decreases image noise and allows you lower radiation dose
Exposure of the entire cardiac cycle
  Either expose the entire cardiac cycle at a constant tube current setting without dose modulation but use a lower tube current than used for coronary CTA
  Or employ dose modulation with a peak window in systole and lowered but diagnostic tube current outside the acquisition window

## Slide 31: CT assessment
Phase selection and measurement tools
Measurements should be obtained at peak systole, where the main pulmonary artery (MPA) has the largest dimensions
Ideally, assessment is performed with a centerline technique, yielding cross-sections of the MPA; contouring tools yield perimeter and min/max diameters

## Slide 32: CT assessment
Center Line Technique / Curved Planar Reformat (CPR)
For ‘parallel/symmetric’ or ‘crisscross’ branch PA origin, align centerline with MPA, beginning in the RVOT and extending to the ridge of bifurcation

## Slide 33: CT assessment
Center Line Technique / Curved Planar Reformat (CPR)
In the setting of highly asymmetric branch PAs with acute take-off of the non-dominant branch PA, extend the centerline into the dominant branch PA

## Slide 34: CT assessment
Annular level
Transition between RVOT and MPA
Identified based on native leaflet tissue (if present), but at times difficult to define
Bifurcation level
Beginning of the first branch MPA take-off
Should not measure too far distally, as cross-sectional dimensions may significantly increase
Annular level
Bifurcation level

## Slide 35: CT assessment
Mid-MPA and additional measurements
Additional cross-sectional assessment is performed at Mid-MPA, halfway in between annular level and bifurcation
If applicable, measure at any level of particular consideration (narrowest, widest)
Bifurcation level
Mid-MPA
Annular level
Bifurcation level
Mid-MPA

## Slide 36: CT assessment
Landing zone length
Distance from the annular plane to beginning of the bifurcation along the centerline
Most post-processing platforms with centerline function display distance between the ‘slider’ positions
Annular level
Bifurcation level

## Slide 37: CT assessment
Distance to ridge of bifurcation
The distance from the beginning of the bifurcation to the ridge is important to estimate how far the stent struts may extend into the bifurcation
Measure from beginning of bifurcation to ridge of bifurcation along centerline
Annular level
Bifurcation level

## Slide 38: CT assessment
RVOT contractility
Can be assessed using cine review of multiphasic data sets
Contractile RVOT
  Muscular tissue contracts in systole and relaxes in diastole
  Results in larger RVOT cross-sectional dimensions in diastole compared to systole
Non-contractile RVOT
  Absence of myocardial tissue (e.g., due to trans-annular patch) may result in lack of diameter change or outward bulging of scar tissue/patch in systole
  Can result in larger RVOT cross-sectional dimensions in systole compared to diastole

## Slide 39: CT assessment
Distance from annular plane to first branch PA
Distance from first branch PA to ridge of MPA bifurcation
Right PA branch diameter – systole
Left PA branch diameter – systole
RVOT/MPA Calcification severity (none, non-protruding, protruding)
RVOT/MPA Calcification extend (focal, horse-shoe, circumferential)
Native pulmonic leaflet present
Coronary anomaly
Coronary artery proximity

## Slide 40: Landing zone assessment
Ensure the zone where the inflow sealing and the outflow portion of the Alterra lands is within sizing parameters
Ensure the ideal location has sufficient length to accommodate the Alterra
  ≥ 48 mm from contractile tissue to back wall MPA
Identify existing structures within RVOT and landing zone (i.e., leaflet remnants, anatomical structures, calcium)
Alterra adaptive prestent sizing | 
Perimeter | 84.9 mm – 119.3 mm
Perimeter derived diameters | 27 mm – 38 mm

## Slide 41: Alterra positioning considerations
Shape of the RVOT/PA
Contractile tissue
Native leaflets
Calcification patterns
In situ devices
  PA stents
Angulation and takeoff of branch pulmonary arteries
NOTE
    Ensure fabric covered portion of prestent does not cover pulmonary branches.

## Slide 42: Alterra screening report
Volume rendering with Alterra

## Slide 43: Stent fracture considerations
The RVOT is a dynamic structure, and implanted prestents can be exposed to complex cyclic stressors
Risk factors for prestent fractures are not fully defined
Potential risk factors may include:
  Anatomical acute angle next to inflow prestent apices
  Engagement of outflow apices with that back wall of the pulmonary artery
  Prestent interaction with other devices during the procedure
Outflow apex engaged with pulmonary roof in bifurcation
Inflow apex meeting acute angle of the outer curvature of the vessel

## Slide 44: Staging considerations
There is the ability to stage procedure if there is concern about stability of the prestent in anatomy to allow for endothelialization
Pre-procedural considerations for staging
Limited acceptable zone for deployment secondary to apex engagement
Borderline large patient anatomy
Procedural considerations for staging
  Inflow apices are not engaged in anatomy
  Excessive motion of the prestent within the anatomy
  Wall apposition is not apparent

## Slide 45: Alterra procedural overview

## Slide 46: Patient preparation
Prepare patient per facility standard practice for pulmonary valve implantation
Vascular access
  Arterial
  5Fr or similar access – aortography, coronary assessment
  Venous
  7Fr or similar access – serial RVOT angiography, pacing (optional)
  Device access site
NOTE
    If IJ access warranted, consideration for location of angiographic catheter placement should be made if femoral venous anatomy is not adequate/sufficient
    A 5Fr or 6Fr angiographic catheter may be put through the same 26Fr sheath if there is only a single point of venous access

## Slide 47: Wire placement
Utilize balloon tip catheter to position wire as far into the pulmonary vascular bed as possible
Consideration of pulmonary artery selection for wire position is critical to facilitate tracking and positioning of the Alterra prestent and the SAPIEN 3 Pulmonic delivery system
Stiff guidewire is recommended to facilitate the tracking of the Alterra prestent and SAPIEN 3 Pulmonic delivery system (e.g., Lunderquist/Meier)
Click above to pause/play video
Click above to pause/play video

## Slide 48: Wire placement – RPA vs LPA
Considerations for wire positioning
Coaxiality with the landing zone
Support for tracking Alterra and the pulmonic delivery system
Click above to pause/play video
Click above to pause/play video
NOTE
    For patients that present without previous surgical intervention it is strongly recommended to utilize RPA wire positioning to assist with coaxial positioning within the MPA.

## Slide 49: Wire placement – RPA vs LPA
RPA wire position attempt
Note the posterior location of the Alterra and the lower position near the STJ on the lateral projection
Device recaptured and removed
Click above to pause/play video
Click above to pause/play video

## Slide 50: Wire placement – RPA vs LPA
Wire positioned into the LPA
Alterra advanced into LPA and deployment initiated
Click above to pause/play video
Click above to pause/play video

## Slide 51: Sheath insertion
The proximal tapered end of the sheath is larger in diameter
Hydrate sheath but do not wipe off hydrophilic coating
Insert the sheath with the Edwards logo facing upward so the seam remains down to ensure proper sheath performance
Insert slowly with continuous motion to minimize friction
Ensure fully expandable portion remains completely within the vessel for proper hemostasis
Ensure the sheath tip is inserted beyond the bifurcation
Partially expandable
Fully expandable
NOTE
    Alterra adaptive prestent is compatible with the 16Fr eSheath or equivalent short sheath. Alterra delivery system outer diameter is 6.7 mm / 20Fr
    Do not use if the sheath is damaged (i.e., kinked or stretched)

## Slide 52: Landing zone interrogation
NOTE
    It is recommended to perform the RV angiogram with the stiff working wire in position.
Baseline right heart pressures
RVOT/PA angiography
  Identification of landmarks
  Determine ideal c-arm angles (AP/CRA, lateral)
  Recommended to utilize both AP/CRA and LAT views to assist with visualization in the 3-dimensional space
Click above to pause/play video
Click above to pause/play video

## Slide 53: Landing zone interrogation
Balloon interrogation for borderline anatomies
Perform with soft balloon i.e., PTS/Amplatzer sizing balloon
Optional baseline ICE imaging at physician preference
Identification of anatomical features
  Leaflet remnants
  Baseline tricuspid valve interrogation
X1 Distance: 23.57 mm
X2 Distance: 26.47 mm
X3 Distance: 26.35 mm

## Slide 54: Delivery system insertion
NOTE
    Insertion force through the partially expandable portion of the sheath can be higher than the push force through the fully expandable portion.
Insert and advance the delivery system
Important: Advance the delivery system from the shaft
  Do not push the delivery system in from the handle
  Do not rotate the wheel during advancement of the delivery system
Click above to pause/play video

## Slide 55: Deployment strategies
Continue to advance the Alterra delivery system until at the desired location for initiating deployment
  “Pull back technique” Initiate deployment distal to intended landing zone the retract the Alterra to the intended landing zone
  “Stationary technique” Initiate deployment at the intended landing zone
Utilize stent connector as inflow positioning marker
Stent connector
Click above to pause/play video

## Slide 56: Deployment strategies
Stepwise deployment pausing with marker band at:
  ~30% - delivery system marker band at crest of distal “wave” of Alterra
  ~50% - delivery system marker band just proximal to waist markers
  ~65% - delivery system marker band at crest of proximal “wave” of Alterra
At the pauses perform angiography to assess:
  Alterra positioning
  Coaxiality of the Alterra within the MPA
  Deployment angulation of outflow apices
~30%
~65%
~50%

## Slide 57: Alterra deployment
NOTE
    The delivery system marker band is located slightly proximal to the distal edge of the outer shaft.
    The prestent can be recaptured into the outer shaft and repositioned if deployed approximately 65%.
While maintaining back tension on the delivery system begin to rotate the deployment wheel towards the open Alterra (counterclockwise) until the Alterra beings to “flower” out of the delivery system
Click above to pause/play video
CAUTION
    Once deployment has begun do not reposition the device more distally. Advancement of the device with the prestent exposed may increase the risk for vascular damage.

## Slide 58: Initial deployment & positioning example – “Pull-back technique”
Previous generation device used for illustration
Click above to pause/play video
Click above to pause/play video
Click above to pause/play video

## Slide 59: Initial deployment & positioning example – “Pull-back technique”
Click above to pause/play video
Click above to pause/play video
Click above to pause/play video

## Slide 60: Initial deployment & positioning example – “Stationary technique”
Click above to pause/play video
Click above to pause/play video
Click above to pause/play video

## Slide 61: Alterra deployment
Once the outflow of the device is in the desired position continue rotating the deployment wheel towards the open Alterra (counterclockwise) until ~65% of the Alterra is deployed
At ~65% deployment, the delivery system marker band will be at crest of the crimped proximal/inflow sealing portion of the Alterra prestent
~65%
CAUTION
    Use of excessive contrast media may lead to renal failure. Measure the patient’s creatinine level prior to the procedure. Contrast media usage should be monitored.

## Slide 62: Alterra recapture
~65%
The Alterra prestent and delivery system is designed with the ability to recapture and reposition the device one time if needed
Once the device is deployed to ~65% assess the device positioning within the landing zone
If repositioning is needed follow these steps:
  Rotate the deployment wheel towards the closed Alterra (clockwise) until the radiopaque marker band is covering the distal apices
  Reposition the delivery system to the intended landing zone
  Redeploy the Alterra prestent
Click above to pause/play video
NOTE
    Several rotations of the deployment wheel may be necessary before the Alterra begins to be recaptured. This is due to the need for the delivery system to overcome the forces to recapture the Alterra prestent
    The prestent can be recaptured and redeployed one (1) time. If a second recapture of the partially deployed prestent is performed, remove and replace the device
CAUTION
    Do not overdrive the outer shaft onto the tapered tip when recapturing the prestent. This may cause the delivery system to cinch down on the guidewire preventing independent movement of the delivery system and guidewire
    Recapturing and redeploying a prestent more than one (1) time may impact implant integrity
    Recapturing a prestent that has been deployed more than 65% may cause system damage

## Slide 63: Alterra deployment
Once the device is at the intended landing zone continue rotating the deployment wheel towards the open Alterra (counterclockwise) until the Alterra prestent is released from the delivery system
    The radiopaque marker band should be beyond the stent connector to ensure release of the prestent
Click above to pause/play video

## Slide 64: Alterra release
Check in multiple orthogonal angles to ensure the prestent connector tabs have released from the prestent connector
If stent connector tabs and/or inflow apicies have not released
  Ensure the outer capsule is proximal to the stent connector
  Attempt to centralize the delivery system
  Catheter and wire manipulations may allow the stent connector tab to release from the stent connector
  Consider changing to softer wire i.e., Amplatz super-stiff
Stent connector tabs still constrained by outer sheath
Stent connector tabs released from stent connector
Radiopaque marker band still over stent connector
Click above to pause/play video

## Slide 65: Alterra release
Previous generation device used for illustration
Outer shaft
Stent connector
Apices still connected
Click above to pause/play video
Click above to pause/play video
CAUTION
    Failure to identity release of the stent connector tab from the stent connector may lead to prestent embolization during removal of the Alterra delivery system.

## Slide 66: Alterra release
Click above to pause/play video
Click above to pause/play video

## Slide 67: Alterra release
Click above to pause/play video
Click above to pause/play video

## Slide 68: Delivery system removal
It is recommended to have the tapered tip in the center of the prestent when retracting
  Ensure the tapered tip and/or the prestent connector does not engage apices
  Wire manipulation may assist in positioning the tapered tip in the center of the prestent
Consider allowing the prestent to fully expand prior to removal of the tapered tip
CAUTION
    Ensure that the tapered tip and delivery system do not interfere with the prestent upon removal to prevent movement of the prestent.
NOTE
    It is not recommended to readvance the outer shaft to recapture the nosecone.

## Slide 69: Delivery system removal
Click above to pause/play video
Click above to pause/play video

## Slide 70: Delivery system removal
If resistance is met upon retraction of the tapered tip:
Stop – ensure the tapered tip is not engaged with the apices
Re-center the tapered tip within the prestent and attempt removal again
If unable to re-center the tapered tip, consider wire exchange with a softer wire
  Ensure the tapered tip is readvanced beyond the open cells prior to exchanging for a softer wire
Click above to pause/play video

## Slide 71: Delivery system removal
Once the tapered tip has cleared the outflow apices continue retracting the entire delivery system into the IVC
Remove the delivery system through the sheath
It is not necessary to recover the delivery system prior to removal
Click above to pause/play video

## Slide 72: Post deployment angiography
Assess stent for stability
Considerations for staging:
  Inflow apices are not engaged in anatomy
  Excessive motion of the prestent within the anatomy
  Wall apposition is not apparent
  Change in Alterra location during delivery system removal
Click above to pause/play video
Click above to pause/play video
CAUTION
    Failure to identity prestent instability may lead to prestent migration/embolization when tracking interventional devices through the prestent.

## Slide 73: Staging considerations
There is the ability to stage procedure if there is concern about stability of the prestent in anatomy to allow for endothelialization
Procedural considerations for staging:
  Inflow apices are not engaged in anatomy
  Excessive motion of the prestent within the anatomy
  Wall apposition is not apparent

## Slide 74: Wire positioning prior to valve insertion
Considerations for changing wire position
  Coaxiality during valve deployment and positioning
  RPA wire position may provide a more coaxial position of the valve within the Alterra
Repositioning wire after Alterra placement
  Critical that a balloon catheter is used to reposition wire
  Ensure that the wire is not through one of the open cells of the Alterra

## Slide 75: Alterra troubleshooting

## Slide 76: Alterra repositioning
If attempting to position the device proximally with gentle retraction on the delivery system and the position of the prestent is not responding
  Relax tension on the delivery system
  Rotate the deployment wheel towards the closed Alterra (clockwise) ½ turn to change the orientation of the outflow apices
  Attempt to reposition the Alterra proximally
If this technique does not work, fully recapture the entire prestent and reposition
Apices outward
Apices after ½ turn rotation

## Slide 77: Ballooning Alterra
Considerations for ballooning Alterra prior to valve insertion
  If Alterra does not appear to be fully expanded ballooning with a soft sizing balloon may be considered
  Consider using a 30 mm or 40 mm PTS or Amplatzer sizing balloon
Insert balloon into unexpanded portion of the Alterra
Slowly inflate and deflate balloon
Previous generation device shown for illustration
Click above to pause/play video
Click above to pause/play video

## Slide 78: Ballooning Alterra
Ballooning performed with 24 mm x 4 cm Vida to 6 ATM
Click above to pause/play video
Click above to pause/play video

## Slide 79: Difficulty retracting Alterra delivery system
If difficulty is experienced when removing/retracting the Alterra delivery system over the wire after recapture:
  Ensure the outer shaft has not been overdriven onto the tapered tip
  Rotate the deployment wheel towards the open Alterra (counterclockwise) to release the tension in the system
  Attempt to remove the Alterra delivery system over the wire

## Slide 80: Difficulty releasing Alterra delivery system
If you are unable to completely retract the outer shaft to release the Alterra:
  Assess the amount of torque and tension you have on the system
  Relax the torque/tension this should allow for the outer shaft to continue to retract on its own
  If the outer shaft continues to not retract:
  The goal is to remove any tension that is on the inner shaft of the delivery system
  Wire manipulation may be used to relieve the inner shaft tension
  Consider changing the wire out to a softer wire i.e., Amplatz super-stiff
Confirm the outer shaft is proximal to the stent connector prior to removing the delivery system

## Slide 81: SAPIEN 3 Pulmonic delivery system procedural overview

## Slide 82: Verify pulmonic THV orientation
Inflow
Outflow
Outflow
Inflow
Verify THV orientation with the inflow (outer skirt) of the THV towards the proximal end (handle) of the delivery system
Valve orientation verification must be performed prior to advancing the valve capsule over the valve
Confirm valve is crimped between shoulders
WARNING
    The physician must verify correct orientation of the valve prior to its implantation.

## Slide 83: Verify correct delivery system volume
Verify the inflation volume with the volume indicated by the printing on the gripper and the color coding
Verify inflation device is locked prior to inserting the delivery system
Colored gripper
Locked
* The delivery system requires a prescribed volume for THV deployment and proper function.
THV Size | Nominal Inflation Volume*
29 mm | 29 mL
CAUTION
    Maintain the inflation device provided by Edwards Lifesciences in the locked position until valve deployment.

## Slide 84: Delivery system and inline sheath insertion
Prior to insertion of the pulmonic valve delivery system:
Ensure tapered tip and valve capsule are hydrated to activate hydrophilic coating
Delivery system is in the locked position
Locked position
Capsule approximated to tapered tip
Ensure that there is no gap between the valve capsule and tapered tip
  If gap is present – unlock the delivery system and retract the balloon catheter until gap is closed
  Relock the delivery system
Remove existing sheath (including Edwards Expandable Introducer Sheath) and prepare venous anatomy with provided dilator
  Consider allowing the dilator to indwell for several minutes prior to insertion of the pulmonic delivery system

## Slide 85: Delivery system and inline sheath insertion
Insert the tapered tip of the delivery system and the inline sheath as a unit until the inline sheath is fully inserted into the vasculature
  Tips for insertion
  Hold the delivery system on back half of capsule
  Ensure the tapered tip and valve capsule remain together until the capsule is inserted in the vasculature
Secure inline sheath
Area for holding during insertion
NOTE
    If the nosecone and capsule become separated, remove the system and re-advance the outer capsule to the nosecone and ensure the delivery system is locked
    Do not force the inline sheath over the valve capsule

## Slide 86: Tracking the THV into the landing zone
Align the Alterra waist markers under fluoro to ensure perpendicular view
Advance the delivery system
Prior to tracking into the heart
  Visualize the valve within the capsule
    Ensure the valve is positioned between the shoulders
    Valve crimp shape is uniform
Maintain the valve capsule over the valve until the valve is positioned at the intended landing zone
Animal model shown for illustration
Click above to pause/play video
NOTE
    Do not force the inline sheath over the valve capsule.

## Slide 87: Tracking the THV into the landing zone
Use care when advancing the pulmonic delivery system through the previously implanted Alterra prestent
  Be aware of any potential gap between the nosecone and the capsule of the delivery system
  Consider slowly reapproximating the nosecone and capsule if a gap is noted. Use care to not overdrive the capsule over the nosecone
Use wire techniques to
  Centralize the wire in the Alterra prior to advancing the pulmonic delivery system
  Position the wire in the lumen of the Alterra and not in the “V” of the apices
Animal model shown for illustration
Click above to pause/play video
CAUTION
    Use caution when advancing devices/delivery systems into the implanted Alterra adaptive prestent to avoid engagement with the inflow apices.

## Slide 88: Positioning the THV within Alterra
60%
40%
Alterra Waist
Place the SAPIEN 3 completely within the waist of the Alterra Prestent
Utilize multiple fluoroscopic views
Alterra waist
Animal model shown for illustration

## Slide 89: Retracting the valve capsule
Once the valve is positioned in the intended landing zone
To uncover the valve
  Unlock the delivery system
  While maintaining the position of the balloon shaft retract the handle until the handle stops at the gripper
    The valve is fully uncovered when the handle meets the gripper
Relock the delivery system
NOTE
    Maintain inline sheath position
    Maintain guidewire position in the pulmonary artery during valve unsheathing to prevent loss of guidewire position

## Slide 90: Verify position of THV
60%
40%
Alterra Waist
Under fluoroscopy verify the position of the THV
  The valve may move during uncovering
If additional positioning is required, ensure the delivery system is locked and maneuver the handle/balloon catheter as a unit
Ensure the handle and the gripper are together and locked
SAPIEN 3 positioning within the waist
  60% SAPIEN 3 valve frame outflow
  40% SAPIEN 3 valve frame inflow
Final implant position of the SAPIEN 3 valve targeted for the middle of the Alterra waist
Locked position
Animal model shown for illustration

## Slide 91: Positioning the THV within Alterra
Positioning considerations
Movement during deployment – valve tends to move towards the ventricle during inflation
Pulmonary branch location in relationship to balloon catheter
  For more distal Alterra deployments, the valve balloon catheter may interact with the branch pulmonary arteries. This tends to push the valve delivery system more ventricular during balloon inflation
Coaxiality of crimped valve within the Alterra visualized in multiple views
Valve fixation during deployment
  The valve will fixate on the first point of interaction with the waist of the Alterra

## Slide 92: Verify position of THV
Under fluoroscopy verify the position of the THV
If additional positioning is required:
  Manipulate the handle and the balloon shaft as a unit
  Ensure the handle and the gripper are touching, and the delivery system is locked
  If unable to reposition the valve:
    Recover the valve with the capsule
    Once the valve is recovered, re-position the valve within the intended landing zone, ensuring the delivery system is locked and the handle and gripper move together as a unit
    Retract the valve capsule as instructed on page 86
Click above to pause/play video
NOTE
    Use care not to entrap the proximal portion of the Alterra frame during recovering of the THV.
Animal model shown for illustration

## Slide 93: Rapid pacing protocol – optional
Protocol for pulmonic THV deployment |  | 
Operator | Sequence | Verbal
Primary operator | THV in position | “Ready?”
Secondary operator | Unlock inflation device | “Ready”
Primary operator | Verify THV properly positioned | “Begin inflating”
Secondary operator | Begin slow, controlled inflation | “Inflating”
Primary operator | Assess THV position | “Hold”
Secondary operator | Stop inflation | “Holding”
Primary operator | Reposition THV as necessary | “Inflate”
Secondary operator | Continue slow, controlled inflation | “Inflating”
Primary operator | Confirm final THV position | “Fully inflate”
Secondary operator | Fully inflate, hold for 3 seconds, and deflate | “ 1… 2… 3… deflating”
Repeat as necessary
NOTE
    Secondary operator should only act on the direction of the primary operator.

## Slide 94: THV deployment
Check to ensure:
Valve capsule is fully retracted prior to deployment
Gripper and the handle are touching
Delivery system is locked
Additional considerations
Verify valve capsule is fully retracted and the handle and gripper are together
  Ensure full inflation of balloon during deployment
  Ensure stability of delivery system during THV deployment
Slow inflation during initial deployment may help with stability of the delivery system and THV during deployment
If considered, reposition only at the very early stage of deployment
Always maintain control of the plunger of inflation device when releasing it
Never lock the inflation device during balloon pre-dilation or THV deployment
  Unlock the inflation device
  Confirm positioning of the valve within intended landing zone
  Begin initial deployment with a slow controlled inflation
  Reposition THV as necessary
  Fully inflate and hold for 3 seconds to ensure complete deployment
  Completely deflate the balloon and withdraw
3 Sec

## Slide 95: THV deployment
SAPIEN 3 THV deployment characteristics
Possible movement during initial deployment
  A slow, controlled inflation during initial deployment may help with stability of the delivery system and THV
Foreshortening visible toward the end of THV deployment
THV Size | Crimped height1 | Expanded height | Foreshortening
29 mm | 31 mm | 22.5 mm | 8.5 mm
1 Rounded to the nearest 0.5
Click above to pause/play video
Animal model shown for illustration

## Slide 96: Withdrawing the SAPIEN 3 Pulmonic delivery system
Once the balloon is fully deflated retract the delivery system into the vena cava
It is recommended to have the tapered tip in the center of the prestent and valve when retracting
  Ensure the tapered tip does not engage outflow apices
  Wire manipulation may assist in positioning the tapered tip in the center of the prestent and valve
If resistance is met upon retraction of the tapered tip:
  Stop – ensure the tapered tip is not engaged with the outflow apices
  Re-center the tapered tip within the prestent and valve and attempt removal again
  If unable to re-center the tapered tip, consider wire exchange
Care should be used to maintain the wire and inline sheath position
Click above to pause/play video
Animal model shown for illustration

## Slide 97: SAPIEN 3 Pulmonic delivery system removal
When the delivery system is in the vena cava (under fluoro), re-sheath the balloon by retracting the gripper until the edge of the valve capsule is touching the tapered tip
Unlock the delivery system lock
Retract the balloon catheter until the end of travel is reached and the balloon shaft marker is exposed
Lock the delivery system
Balloon shaft marker
Animal model shown for illustration
Predicate device shown for illustrative purposes
Click above to pause/play video
CAUTION
    Completely cover the balloon prior to removal to minimize the risk of vascular injury.

## Slide 98: SAPIEN 3 Pulmonic delivery system removal
3
Retract the balloon catheter until the end of travel is reached and the balloon shaft marker is exposed
5
Lock the delivery system
1
Inflation device unlocked
4
If unable to completely retract the balloon shaft
  Advance the balloon shaft
  Rotate the balloon
  Retract the balloon shaft into the valve capsule
2
Unlock the delivery system lock
Click above to pause/play video

## Slide 99: Difficulty retracting balloon into valve capsule
If difficulty is experienced when attempting to recover the balloon catheter:
Neutral pressure technique
Balloon rotation technique
  Advance the balloon shaft
  Rotate the balloon
  Retract the balloon shaft into the valve capsule
  Advance the balloon shaft
  Place the inflation device at neutral pressure
  Retract the balloon shaft into the valve capsule
Balloon re-folding technique
  Advance the balloon shaft
  Inflate with a small amount of volume
  Deflate the balloon
  Retract the balloon shaft into the valve capsule

## Slide 100: SAPIEN 3 Pulmonic delivery system removal
Continue to retract the delivery system until the capsule meets the inline sheath tip
Retract the entire system (inline sheath and delivery system) together out of the vasculature
Replace system with standard vascular access sheath, e.g., Gore DrySeal or other large bore sheath per physician preference
THV Size | Catheter capsule outer diameter
29 mm | 28.5Fr / 9.5 mm

## Slide 101: Post deployment assessment
Assess the valve and prestent post deployment using fluoroscopy
Insert angiography catheter and remove wire
Perform angiography and hemodynamics
Perform post angiogram to assess:
  Valve position and expansion within the Alterra prestent
  Functional assessment of prestent and valve (central leak, para-stent leak, PVL, etc.)
Close vasculature per physician preference
Click above to pause/play video
Click above to pause/play video
NOTE
    Stiff wire tends to exacerbate central regurgitation.

## Slide 102: Post-dilation (if necessary)
Consider post-dilation if paravalvular jets are clinically significant
Post-dilate with the pulmonic delivery system if not previously removed or other commercially available balloon catheter
If using the pulmonic delivery system, utilize the radiopaque shoulder markers to position the balloon within the THV prior to inflation
Shoulder markers
NOTE
    If regurgitation is central rather than paravalvular, DO NOT post-dilate.

## Slide 103: Intra-procedural considerations

## Slide 104: Removing undeployed delivery system
Prior to removal of the delivery system from the anatomy the valve must be covered by the capsule
Considerations for location to recover the valve
  Valve location within the Alterra
  Stability of the Alterra
  Valve and capsule location anatomically
    Position of the capsule in respect to the tricuspid valve and surrounding anatomy
Ensure the delivery system freely moves upon retraction into the vena cava
Continue to retract the delivery system until the capsule meets the inline sheath tip
Retract the entire system (inline sheath and delivery system) together out of the vasculature
NOTE
    Equipment will need to be ready to insert into vasculature prior to removal of the pulmonic valve delivery system.
CAUTION
    Completely cover the balloon prior to removal to minimize the risk of vascular injury.

## Slide 105: Complications

## Slide 106: Balloon rupture
If delivery system balloon ruptures or leaks during deployment without THV embolization
Do not use excessive force take care when crossing the THV and tracking through the anatomy
Based on an assessment of the size and tortuosity of the pulmonary artery anatomy, evaluate the risks and tradeoffs of the following delivery system retrieval options
Maintain guidewire position
Check for paravalvular or central leaks utilizing echo
If post-dilation needed, use a new delivery system or other commercially available balloon catheter
Option 1
  Utilize snare technique (must be performed prior to attempting retrieval into the capsule)
    If the tear is radial, it may umbrella and make retrieval into the capsule extremely difficult
    Utilization of the snare prior to attempting to retrieve a radial tear may reduce the profile of the balloon for removal
Option 2
  Retract the delivery system into the vena cava
    Once in the vena cava and retrieve the delivery system (through the tip of the sheath)
    Attempt to snare the ruptured balloon in the vena cava prior to retraction into the capsule

## Slide 107: Balloon rupture: Snaring in pulmonary artery
Balloon should remain in the pulmonary artery if the patient is hemodynamically stable
Retract guidewire until it is flush with nosecone but no further. Position the snare distal to nosecone
Maneuver the opened snare over the distal end of the delivery system. Gently tighten snare while over the distal end of the balloon (approximately the level of the distal balloon shoulder marker)
Retract the delivery system and snare into the vena cava
Retract the delivery system while advancing catheter/snare until the delivery system is in the femoral
3
5
1
2
3
4
5
6
Access the distal PA via the contralateral femoral. Insert single loop snare
7
Once in femoral, attempt full recapture of the balloon into the capsule or convert to vascular surgery for device removal
WARNING
    Do not apply excessive force if any resistance is encountered.
NOTE
    After snaring balloon, guidewire position may be adjusted outside of delivery system to maintain wire access
    Do not allow snare to drag across femoral bifurcation during removal

## Slide 108: Pulmonary perforation
Risk factors
Overview
  Rare but serious complication that is avoidable with good wire practices
  Stiff wires
  Wire manipulation
  Wire replacement
Symptoms & diagnostics
Prevention overview
  Will see contrast extravasation
  Large perforations may result in hemodynamic changes
  Wire control, use the softest wire possible, observe wire position at all times
  If perforation occurs, treatment depends on location and size of perforation
    Small: Watch and wait
    Large: Covered stent, vascular occlusion devices, balloon tamponade

## Slide 109: Embolization of device
  Maintain guidewire position through the device
Assess patient hemodynamic stability
If embolization into the ventricle, consider CPB and surgical intervention
If embolization further into the pulmonic artery assess stability – if needed, consider CPB and surgical intervention

## Slide 110: Alterra or valve malposition
Maintain wire position through the device
Assess patient hemodynamic stability
Consider percutaneous or surgical intervention
Previous generation device shown for illustrative purposes
Click above to pause/play video

## Slide 111
CAUTION: Federal (United States) law restricts these devices to sale by or on the order of a physician. See Instructions for Use for full prescribing information, including indications, contraindications, warnings, precautions, and adverse events.
Edwards, Edwards Lifesciences, the stylized E logo, Alterra, Edwards SAPIEN, Edwards SAPIEN 3, Edwards eSheath, eSheath, Qualcrimp, SAPIEN and SAPIEN 3 are trademarks of Edwards Lifesciences Corporation. All other trademarks are the property of their respective owners.
© 2025 Edwards Lifesciences Corporation. All rights reserved.