## Slide 1: Edwards SAPIEN 3 Transcatheter Pulmonic Valve Delivery System
20 / 23 / 26 / 29 mm Sizes
Pulmonic Procedural Training Manual
See instructions for use for full prescribing information, including indications, contraindications, warnings, precautions and potential adverse events
Edwards Confidential. Do not reproduce or distribute.
DOC-0250104 Rev A

## Slide 2: Table of contents
System overview
Procedural animation
04
Edwards SAPIEN 3 pulmonic delivery system overview
05
Device specifications
Edwards SAPIEN 3 THV
08
Edwards SAPIEN 3 pulmonic delivery system
11
Recommended equipment
20
RVOT angiogram
22
Procedural steps
Wire exchange and placement
23
Coronary occlusion assessment
24
Prestenting considerations
30
Landing zone sizing
29
Pulmonic SAPIEN 3 THV sizing
31
Verify pulmonic THV orientation
33
35
Delivery system and inline sheath insertion

## Slide 3: Table of contents
Procedural steps (cont’d)
Verify position of THV
41
Rapid pacing protocol - optional
42
Post-dilation (if necessary)
45
THV deployment
43
Withdrawing the SAPIEN 3 Pulmonic delivery system
46
SAPIEN 3 Pulmonic delivery system removal
47
51
Post-deployment THV assessment
Intraprocedural considerations
53
Removing undeployed delivery system
54
Balloon rupture
39
THV positioning
40
Retracting the valve capsule
Tracking the THV into landing zone
37
Procedural steps (cont’d)

## Slide 4: Edwards SAPIEN 3 Pulmonic System
Procedural animation of the THV pulmonic procedure
Click above to pause/play video

## Slide 5: Edwards SAPIEN 3 pulmonic delivery system overview

## Slide 6: SAPIEN 3 Pulmonic system with the SAPIEN 3 Transcatheter heart valve
Edwards SAPIEN 3 Transcatheter heart valve
Crimper
Pulmonic delivery system
Inflation devices
Accessories
Color-coded system
Pink
20 mm THV
Green
23 mm THV
Purple
26 mm THV
Orange
29 mm THV

## Slide 7: Device specifications

## Slide 8: Edwards SAPIEN 3 THV
Outer skirt
PET outer skirt designed to minimize paravalvular leak
Bovine pericardial tissue
Scalloped leaflet shape
Carpentier-Edwards ThermaFix* process is intended to minimize the risk of calcification
Inner skirt
Polyethylene terephthalate (PET)
* No clinical data are available which evaluate the long-term impact of the Carpentier-Edwards ThermaFix process in patients.

## Slide 9: Edwards SAPIEN 3 THV
Frame design
Wide strut angles contribute to ultra-low delivery profile
Four rows and columns between commissures for high radial strength
Cobalt-chromium alloy frame and wide strut angles provide fatigue resistance and high radial strength for circularity
Chart: 
| 23 mm | 26 mm | 29 mm
SAPIEN XT | 1.26 | 1.08 | 1.19
SAPIEN 3 | 1.44 | 1.37 | 1.3
Avg radial strength (bar)
Edwards Lifesciences internal data: 2-sample t-test to determine avg radial strength of Edwards frames for SAPIEN 3 > SAPIEN XT, P <.05

## Slide 10: Edwards SAPIEN 3 THV
1 No clinical data are available which evaluate the long-term impact of the Edwards Lifesciences tissue treatment in patients.
2 Rounded to the nearest 0.5 mm
3 Measured at bottom of zigzag
 | 20 mm | 23 mm | 26 mm | 29 mm
Tissue | Bovine pericardium, ThermaFix tissue treatment1 |  |  | 
Frame | Cobalt-chromium alloy |  |  | 
Noncompliant landing zone size during balloon predilation | 16.5 - 20 mm | 20 - 23 mm | 23 - 26 mm | 26 - 29 mm
Crimped height2 | 21 mm | 24.5 mm | 27 mm | 31 mm
Expanded height | 15.5 mm | 18 mm | 20 mm | 22.5 mm
Foreshortening | 5.5 mm | 6.5 mm | 7 mm | 8.5 mm
Inner skirt height3 | 7.9 mm | 9.3 mm | 10.2 mm | 11.6 mm
Outer skirt height | 5.2 mm | 6.6 mm | 7.0 mm | 8.1 mm

## Slide 11: SAPIEN 3 Pulmonic delivery system overview
Balloon folding technology simplifies the procedure
Eliminates valve alignment and flex catheter retraction steps
Designed for valve deliverability
Valve capsule to cover SAPIEN 3 valve
Hydrophilic coating on tapered tip and valve capsule
Valve crimped directly on balloon
Inline sheath
Color coded gripper
Tapered tip to facilitate tracking
Valve capsule to cover valve

## Slide 12: SAPIEN 3 Pulmonic delivery system – Specifications
THV size | Catheter capsule outer diameter
20 mm | 25.1Fr / 8.4 mm
23 mm | 27.1Fr / 9.0 mm
26 mm | 27.1Fr / 9.0 mm
29 mm | 28.5Fr / 9.5 mm
112 cm
Effective length balloon uncovered (all sizes)
42 mm Tapered tip (all sizes)
Pulmonic delivery system
Guidewire compatibility
0.035”

## Slide 13: Edwards Pulmonic delivery system – Balloon
Balloon dimensions | 20 mm | 23 mm | 26 mm | 29 mm
Balloon length shoulder-to-shoulder (A) | 24 mm | 28 mm | 30 mm | 34 mm
Proximal shoulder-to-nosecone (B) | 8.0 cm | 8.6 cm | 9.1 cm | 9.5 cm
A
B
THV Size | Nominal inflation volumes* | Rated burst pressures
20 mm | 10 mL | 9 atm
23 mm | 15 mL | 9 atm
26 mm | 21 mL | 9 atm
29 mm | 29 mL | 9 atm
* The delivery system requires a prescribed volume for THV deployment and proper function

## Slide 14: SAPIEN 3 Pulmonic delivery system – Handle
Color-coded gripper
Delivery system lock
Inflation volume
Outer shaft
Flush port
Balloon shaft
Balloon inflation port
Guidewire lumen
Balloon shaft marker

## Slide 15: SAPIEN 3 Pulmonic delivery system – Distal end
Valve is covered by capsule until valve is positioned in desired location
Valve is crimped directly on balloon (no valve alignment)
Animal model shown for illustration
Radiopaque shoulder markers
Tapered tip
Valve capsule
Internal shoulders (underneath balloon)
Valve capsule
Tapered tip
Balloon shaft

## Slide 16: SAPIEN 3 Pulmonic delivery system – Inline sheath
10 cm effective length
25.5Fr / 8.5 mm outer diameter (all sizes)
Flush port
NOTE
    The inline sheath does not come off the delivery system. Shown separately for illustrative purposes.

## Slide 17: SAPIEN 3 Pulmonic delivery system – Dilators
40 cm all sizes
THV size | Provided dilator size
20 mm | 25Fr
23 mm | 27Fr
26 mm | 27Fr
29 mm | 28Fr

## Slide 18: Crimper, Qualcrimp, and 2-piece crimp stopper
Additional components used during device preparation
  Crimper is the same for all THV sizes and is packaged separately
  Qualcrimp crimping accessory is designed to protect the THV leaflets when crimping to a low profile and is packaged with the delivery system
  Color-coded 2-piece crimp stopper is used during crimping and is packaged with the delivery system
Qualcrimp crimping accessory
Crimper
2-piece crimp stopper
23 mm
26 mm
29 mm
20 mm

## Slide 19: Inflation devices
For use with 20 / 23 / 26 mm
THV systems
For use with
29 mm THV system
Inflation devices are used for deployment of the THV and predilation balloon
Volume-based inflation
Lock / unlock symbols on handle
1 mL graduation marks for volume accuracy
Locking Syringe (38 mL)
Used with the 29 mm THV system
Customized long narrow barrel for inflation with large volumes
Pressure gauge removed for device placement during inflation
Inflation device (25 mL)
Used with the 20 mm, 23 mm, and 26 mm THV systems

## Slide 20: Recommended equipment
Sterile table equipment
6 sterile rinsing basins
  2 – physiological saline (1 liter) – valve rinsing
  2 – physiological saline (1 liter) – Qualcrimp rinsing
  1 – heparinized saline (1 liter) – flushing
  1 – diluted radiopaque contrast medium (15 cc contrast: 85 cc saline)
20 cc syringes or larger
50 cc syringes or larger
High-pressure 3-way stopcock
Prep table
Sterile forceps
Sterile scissors
Room equipment
Vascular dilators
Balloon tip catheter
Angiographic catheters
Exchange length 0.035" (0.89 mm) extra-stiff guidewires: i.e., Lunderquist, Meier, Amplatz extra stiff
Intra-cardiac echocardiographic system (optional)
Vascular closure devices
Large bore short sheath
Temporary pacemaker and pacing leads
Liter heparinized saline
Liter normal 0.9% saline
Standard cardiac catheterization lab equipment
Fluoroscopy (bi-plane, fixed, mobile or semi-mobile fluoroscopy systems appropriate for use in percutaneous coronary interventions)
Equipment provided by Edwards Lifesciences
Edwards SAPIEN 3 valve
Edwards Pulmonic delivery system
Edwards crimper
Edwards provided inflation device

## Slide 21: Procedural steps

## Slide 22: RVOT angiography
Click above to pause/play video
Biplane angiography
Assess proposed site for implantation
Quantification of pulmonary regurgitation and/or stenosis
Click above to pause/play video

## Slide 23: Wire exchange and placement
Wire positioned in a distal pulmonary branch
  Preferably the left lower branch given its vertical orientation
Exchange to a 0.035” exchange length stiff guidewire
  i.e., Lunderquist (Cook) or Meier (Boston Scientific)
  Stiff wires are recommended to assist with tracking of the delivery system

## Slide 24: Coronary occlusion assessment
Rotational and/or biplane coronary angiography (typically lateral and LAO / CAU) profiles the left coronary well
Balloon sized to the final anticipated size of landing zone
Ensure the presence of at least 10 mm from the edge of the inflated balloon to the origin of the coronary arteries
Either coronary can be compromised; although typically, it is the left main and/or left anterior descending
Image: Wilson et al. Evolving trends in interventional cardiology: endovascular options for congenital disease in adults
Image: Morray et al. Risk of Coronary Artery Compression Among Patients Referred for Transcatheter Pulmonary Valve Implantation: A Multicenter Experience

## Slide 25: Coronary occlusion assessment
Click above to pause/play video
Example 1
Nonselective technique

## Slide 26: Coronary occlusion assessment
Click above to pause/play video
Example 2
Selective left coronary angiography

## Slide 27: Coronary occlusion assessment
Assess coronaries utilizing multiple angles
Click above to pause/play video
Click above to pause/play video

## Slide 28: Coronary occlusion assessment
Check for remote coronary artery compression

## Slide 29: Landing zone sizing
Measured in multiple planes
Use compliant balloon catheter
Measured at the minimum diameter of the landing zone
Ensure final landing zone diameter doesn’t impinge on coronaries
Ensure a noncompliant landing zone
Eliminate any residual stenosis using standard techniques prior to placing the prosthesis
NOTE
    To minimize the risk of conduit rupture, use caution when using a balloon with a diameter greater than the nominal diameter (original implant size) of the conduit for predilation of the intended deployment site.

## Slide 30: Prestenting considerations
Provides a landing zone for valve and treats the underlying stenosis, especially lesions longer than the valve frame
Helps maintain circular configuration
Predicts final sizing
Bare metal or covered
Pressure measurements post-stent should reflect minimal to no residual gradient
Click above to pause/play video

## Slide 31: Pulmonic SAPIEN 3 THV sizing
Recommended sizing is based on the diameters measured in multiple views of the landing zone
23 mm THV
PULMONIC
20 mm THV
PULMONIC
26 mm THV
PULMONIC
29 mm THV
PULMONIC
Chart: 
| Category 1
Series 1 | 16
Diameter (mm)
THV size | Noncompliant landing zone size during balloon predilation
20 mm | 16.5 - 20 mm
23 mm | 20 - 23 mm
26 mm | 23 - 26 mm
29 mm | 26 - 29 mm

## Slide 32: Pulmonic SAPIEN 3 THV sizing
In presence of presenting, prosthesis sizing should be based on the diameter of the balloon at full inflation

## Slide 33: Verify pulmonic THV orientation
Inflow
Outflow
Outflow
Inflow
Valve orientation verification must be performed prior to advancing the valve capsule over the valve
Confirm valve is crimped between shoulders
WARNING
    The physician must verify correct orientation of the valve prior to its implantation.
Verify THV orientation with the inflow (outer skirt) of the THV towards the proximal end (handle) of the delivery system

## Slide 34: Verify correct delivery system volume
Verify the inflation volume with the volume indicated by the printing on the gripper and the color coding
Verify inflation device is locked prior to inserting the delivery system
Colored gripper
Locked
* The delivery system requires a prescribed volume for THV deployment and proper function
THV Size | Nominal Inflation Volume*
20 mm | 10 mL
23 mm | 15 mL
26 mm | 21 mL
29 mm | 29 mL
CAUTION
    Maintain the inflation device provided by Edwards Lifesciences in the locked position until valve deployment.

## Slide 35: Delivery system and inline sheath insertion
Prior to insertion of the pulmonic valve delivery system:
Ensure tapered tip and valve capsule are hydrated to activate hydrophilic coating
Delivery system is in the locked position
Locked position
Capsule approximated to tapered tip
Ensure that there is no gap between the valve capsule and tapered tip
  If gap is present – unlock the delivery system and retract the balloon catheter until gap is closed
  Relock the delivery system
Remove existing sheath (including Edwards Expandable Introducer Sheath) and prepare venous anatomy with provided dilator
  Consider allowing the dilator to indwell for several minutes prior to insertion of the pulmonic delivery system

## Slide 36: Delivery system and inline sheath insertion
Insert the tapered tip of the delivery system and the inline sheath as a unit until the inline sheath is fully inserted into the vasculature
  Tips for insertion
  Hold the delivery system on back half of capsule
  Ensure the tapered tip and valve capsule remain together until the capsule is inserted in the vasculature
Secure inline sheath
Area for holding during insertion
NOTE
    If the nosecone and capsule become separated, remove the system and re-advance the outer capsule to the nosecone and ensure the delivery system is locked
    Do not force the inline sheath over the valve capsule

## Slide 37: Tracking the THV into the landing zone
Advance the delivery system
Prior to tracking into the heart
  Visualize the valve within the capsule
    Ensure the valve is positioned between the shoulders
    Valve crimp shape is uniform
Maintain the valve capsule over the valve until the valve is positioned at the intended landing zone
Animal model shown for illustration
Click above to pause/play video
NOTE
    Do not force the inline sheath over the valve capsule.

## Slide 38: Tracking the THV into the landing zone
Use care when advancing the pulmonic delivery system to the intended landing zone
  Be aware of any potential gaps between the nosecone and the capsule of the delivery system
  Consider slowly reapproximating the nosecone and capsule if a gap is noted
Use wire techniques to
  Centralize the wire in the RVOT/MPA prior to advancing the pulmonic delivery system
Click above to pause/play video
Animal model shown for illustration
Alterra implant shown for illustrative purposes

## Slide 39: THV positioning
Continue advancing the pulmonic delivery system and position the THV completely within the landing zone
Avoid positioning the distal (outflow) half of the THV within any residual stenosis which may affect proper leaflet function
Use different fluoroscopic views, if uncertain about positioning
Animal model shown for illustration.
Alterra implant shown for illustrative purposes.
Click above to pause/play video
NOTE
    Placement of the THV in the proximity of stenosis can result in THV movement during deployment.

## Slide 40: Retracting the valve capsule
Once the valve is positioned in the intended landing zone
Step 1
Unlock the delivery system
While maintaining the position of the balloon shaft retract the handle until the handle stops at the gripper
  The valve is fully uncovered when the handle meets the gripper
Step 2
Relock the delivery system
Step 3
To uncover the valve:
NOTE
    Maintain inline sheath position
    Maintain guidewire position in the pulmonary artery during valve unsheathing to prevent loss of guidewire position

## Slide 41: Verify position of THV
Under fluoroscopy verify the position of the THV within the intended landing zone
  The valve may move during uncovering
If additional positioning is required, ensure the delivery system is locked and maneuver the handle/balloon catheter as a unit
Ensure the handle and the gripper are together and locked prior to inflation
If unable to reposition the uncovered valve:
  Re-cover the valve with the capsule
  Once the valve is recovered, lock the handle, re-position the valve within the intended landing zone
  Retract the valve capsule as instructed on slide 40
Animal model shown for illustration
Alterra implant shown for illustrative purposes
Locked position
Click above to pause/play video

## Slide 42: Rapid pacing protocol – Optional
Protocol for pulmonic THV deployment |  | 
Operator | Sequence | Verbal
Primary operator | THV in position | “Ready?”
Secondary operator | Unlock inflation device | “Ready”
Primary operator | Verify THV properly positioned | “Begin inflating”
Secondary operator | Begin slow, controlled inflation | “Inflating”
Primary operator | Assess THV position | “Hold”
Secondary operator | Stop inflation | “Holding”
Primary operator | Reposition THV as necessary | “Inflate”
Secondary operator | Continue slow, controlled inflation | “Inflating”
Primary operator | Confirm final THV position | “Fully inflate”
Secondary operator | Fully inflate, hold for 3 seconds, and deflate | “ 1… 2… 3… deflating”
Repeat as necessary
NOTE
    Secondary operator should only act on the direction of the primary operator.

## Slide 43: THV deployment
Check to ensure:
Valve capsule is fully retracted prior to deployment
Gripper and the handle are touching
Delivery system is locked
  Unlock the inflation device
  Confirm positioning of the valve within intended landing zone
  Begin initial deployment with a slow controlled inflation
  Reposition THV as necessary
  Fully inflate and hold for 3 seconds to ensure complete deployment
  Completely deflate the balloon and withdraw
3 Sec
Additional considerations
Verify valve capsule is fully retracted and the handle and gripper are together
  Ensure full inflation of balloon during deployment
  Ensure stability of delivery system during THV deployment
Slow inflation during initial deployment may help with stability of the delivery system and THV during deployment
If considered, reposition only at the very early stage of deployment
Always maintain control of the plunger of inflation device when releasing it
Never lock the inflation device during balloon pre-dilation or THV deployment

## Slide 44: THV deployment
SAPIEN 3 THV deployment characteristics
Possible movement during initial deployment
  A slow, controlled inflation during initial deployment may help with stability of the delivery system and THV
Foreshortening visible toward the end of THV deployment
THV Size | Crimped height1 | Expanded height | Foreshortening
20 mm | 21 mm | 15.5 mm | 5.5 mm
23 mm | 24.5 mm | 18 mm | 6.5 mm
26 mm | 27 mm | 20 mm | 7 mm
29 mm | 31 mm | 22.5 mm | 8.5 mm
1 Rounded to the nearest 0.5
Animal model shown for illustration
Alterra implant shown for illustrative purposes
Click above to pause/play video

## Slide 45: Post-dilation (if necessary)
Consider post-dilation if paravalvular jets are clinically significant
Post-dilate with the pulmonic delivery system if not previously removed or other commercially available balloon catheter
If using the pulmonic delivery system, utilize the radiopaque shoulder markers to position the balloon within the THV prior to inflation
Shoulder markers
NOTE
    If regurgitation is central rather than paravalvular, DO NOT post-dilate.

## Slide 46: Withdrawing the SAPIEN 3 Pulmonic delivery system
Once the balloon is fully deflated retract the delivery system into the vena cava
It is recommended to have the tapered tip in the center of the valve when retracting
  Ensure the tapered tip does not engage previously implanted devices
  Wire manipulation may assist in positioning the tapered tip in the center of the valve
If resistance is met upon retraction of the tapered tip:
  Stop – ensure the tapered tip is not engaged with any in situ devices
  Re-center the tapered tip and attempt removal again
  If unable to re-center the tapered tip, consider wire exchange for a softer wire
Care should be used to maintain the wire and inline sheath position
Animal model shown for illustration
Predicate devices shown for illustrative purposes
Click above to pause/play video

## Slide 47: SAPIEN 3 Pulmonic delivery system removal
When the delivery system is in the vena cava (under fluoro), re-sheath the balloon by retracting the gripper until the edge of the valve capsule is touching the tapered tip
Unlock the delivery system lock
Retract the balloon catheter until the end of travel is reached and the balloon shaft marker is exposed
Lock the delivery system
Balloon shaft marker
Animal model shown for illustration
Predicate devices shown for illustrative purposes
Click above to pause/play video
CAUTION
    Completely cover the balloon prior to removal to minimize the risk of vascular injury.

## Slide 48: SAPIEN 3 Pulmonic delivery system removal
3
Retract the balloon catheter until the end of travel is reached and the balloon shaft marker is exposed
5
Lock the delivery system
1
Inflation device unlocked
4
If unable to completely retract the balloon shaft
  Advance the balloon shaft
  Rotate the balloon
  Retract the balloon shaft into the valve capsule
2
Unlock the delivery system lock
Click above to pause/play video

## Slide 49: Difficulty retracting balloon into valve capsule
If difficulty is experienced when attempting to recover the balloon catheter
Neutral pressure technique
Balloon rotation technique
  Advance the balloon shaft
  Rotate the balloon
  Retract the balloon shaft into the valve capsule
  Advance the balloon shaft
  Place the inflation device at neutral pressure
  Retract the balloon shaft into the valve capsule
Balloon re-folding technique
  Advance the balloon shaft
  Inflate with a small amount of volume
  Deflate the balloon
  Retract the balloon shaft into the valve capsule

## Slide 50: SAPIEN 3 Pulmonic delivery system removal
Continue to retract the delivery system until the capsule meets the inline sheath tip
Retract the entire system (inline sheath and delivery system) together out of the vasculature
Replace system with standard vascular access sheath, e.g., Gore DrySeal or other large bore sheath per physician preference
THV size | Catheter capsule outer diameter
20 mm | 25.1Fr / 8.4 mm
23 mm | 27.1Fr / 9.0 mm
26 mm | 27.1Fr / 9.0 mm
29 mm | 28.5Fr / 9.5 mm

## Slide 51: Post-deployment THV assessment
CAUTION
    Use of excessive contrast media may lead to renal failure. Measure the patient’s creatinine level prior to the procedure. Contrast media usage should be monitored.
Withdraw the delivery system from the valve before assessing
Perform post-angiogram with wire still in position to assess
  Valve position and expansion
  Patency of coronary arteries
  Functional assessment of valve (central leak, PVL, etc.)
Perform pressure assessment
Close vasculature per physician preference
Click above to pause/play video
Assess the valve post-deployment using fluoroscopy
NOTE
    Stiff wire tends to exacerbate central regurgitation.

## Slide 52: Intraprocedural considerations

## Slide 53: Removing undeployed delivery system
Prior to removal of the delivery system from the anatomy the valve must be covered by the capsule
Considerations for location to recover the valve
  Valve location within the pulmonary artery
  Stability of the any previously implanted stents/devices
  Valve and capsule location anatomically
    Position of the capsule in respect to the tricuspid valve and surrounding anatomy
Ensure the delivery system freely moves upon retraction into the vena cava
Continue to retract the delivery system until the capsule meets the inline sheath tip
Retract the entire system (inline sheath and delivery system) together out of the vasculature
CAUTION
    Completely cover the balloon prior to removal to minimize the risk of vascular injury.
NOTE
    Equipment will need to be ready to insert into vasculature prior to removal of the pulmonic valve delivery system.

## Slide 54: Balloon rupture
If delivery system balloon ruptures or leaks during deployment without THV embolization
Do not use excessive force take care when crossing the THV and tracking through the anatomy
Based on an assessment of the size and tortuosity of the pulmonary artery anatomy, evaluate the risks and tradeoffs of the following delivery system retrieval options
Maintain guidewire position
Check for paravalvular or central leaks utilizing echo
If post-dilation needed, use a new delivery system or other commercially available balloon catheter
Option 1
  Utilize snare technique (must be performed prior to attempting retrieval into the capsule)
    If the tear is radial, it may umbrella and make retrieval into the capsule extremely difficult
    Utilization of the snare prior to attempting to retrieve a radial tear may reduce the profile of the balloon for removal
Option 2
  Retract the delivery system into the vena cava
    Once in the vena cava and retrieve the delivery system (through the tip of the sheath)
    Attempt to snare the ruptured balloon in the vena cava prior to retraction into the capsule

## Slide 55: Balloon rupture: Snaring in pulmonary artery
Balloon should remain in the pulmonary artery if the patient is hemodynamically stable.
Retract guidewire until it is flush with nosecone but no further. Position the snare distal to nosecone
Maneuver the opened snare over the distal end of the delivery system. Gently tighten snare while over the distal end of the balloon (approximately the level of the distal balloon shoulder marker)
Retract the delivery system and snare into the vena cava.
Retract the delivery system while advancing catheter/snare until the delivery system is in the femoral.
3
5
1
2
3
4
5
6
Access the distal PA via the contralateral femoral. Insert single loop snare.
7
Once in femoral, attempt full recapture of the balloon into the capsule or convert to vascular surgery for device removal.
WARNING
    Do not apply excessive force if any resistance is encountered.
NOTE
    After snaring balloon, guidewire position may be adjusted outside of delivery system to maintain wire access.
    Do not allow snare to drag across femoral bifurcation during removal

## Slide 56: Pulmonary perforation
Overview
  Rare but serious complication that is avoidable with good wire practices
Risk factors
  Stiff wires
  Wire manipulation
  Wire replacement
Symptoms & diagnostics
  Will see contrast extravasation
  Large perforations may result in hemodynamic changes
Prevention overview
  Wire control, use the softest wire possible, observe wire position at all times
  If perforation occurs, treatment depends on location and size of perforation
    Small: Watch and wait
    Large: Covered stent, vascular occlusion devices, balloon tamponade

## Slide 57: Embolization of device
  Maintain guidewire position through the device
Assess patient hemodynamic stability
If embolization into the ventricle, consider CPB and surgical intervention
If embolization further into the pulmonic artery assess stability – if needed, consider CPB and surgical intervention

## Slide 58
CAUTION: Federal (United States) law restricts these devices to sale by or on the order of a physician. See Instructions for Use for full prescribing information, including indications, contraindications, warnings, precautions, and adverse events.
The clinicians demonstrating Edwards Lifesciences technology in this training material may have received compensation and/or travel reimbursement from Edwards in exchange for providing professional services to Edwards.
Edwards, Edwards Lifesciences, the stylized E logo, Carpentier-Edwards, Edwards Commander, Edwards SAPIEN, Edwards SAPIEN XT, Edwards SAPIEN 3, eSheath, eSheath+, Qualcrimp, SAPIEN, SAPIEN XT, SAPIEN 3, and ThermaFix are trademarks or service marks of Edwards Lifesciences Corporation. All other trademarks or service marks are the property of their respective owners.
© 2025 Edwards Lifesciences Corporation. All rights reserved.